## Revision History
SASP(TSA8C1) Final Test Development Instruction
Rev0.4
CONTENTS
･Measurement system
･Command Format
･control flow and sequence
･test item and command
･Log output format
･Test spec
･BLE measurement item for MT8872
Revision History
Revision | Date | Comments
Rev0.1 | 2021-10-28 | First release
Rev0.11 | 2021-11-09 | Change the title of this file
Rev0.2 | 2021-12-06 | Supports DUT FW Ver1.0
Changed the request command of "End DTM" Before change: 0x01 04 After change: 0xFF FF See "Command Format" sheet. See item No2-12,3-31 on the "test item and command" sheet.
Command change "Set DTM (0x03)" Power mode selection added LDO mode: 0x00 DCDC mode: 0x01 See "Command Format" sheet. See item No2-2,3-2 on the "test item and command" sheet.
Command change "Check 32.768kHz Xtal FREQ (0x08)" Added Enable flag Output disable: 0x00 Output enable: 0x01 See "Command Format" sheet. See item No4-2 on the "test item and command" sheet.
Command change "SET DCDC (0x09)" Added Enable flag DCDC disable: 0x00 DCDC enable: 0x01 See "Command Format" sheet. See item No4-3 on the "test item and command" sheet.
Command addition "Init settings (0x0A)" See "Command Format" sheet. See item No4-3 on the "test item and command" sheet.
Addition of receive command in case of error (0xFF) See "Command Format" sheet.
Rev0.3 | 2022-02-25 | Addition of In-band Spurious(1M/2M). Addition of Out-band Spurious(1M/2M). See item No3-21,3-22,3-23,3-24 on the "test item and command" sheet.
Added Spectrum analyzer and SPDT. See "Measurement system" sheet.
Rev0.4 | 2022-03-04 | In-band Spurious(1M/2M) : command addition. See item No3-21,3-22,3-23,3-24,3-25,3-26 on the "test item and command" sheet. (item number modification) Out-band Spurious(1M/2M) : command addition. See item No3-27,3-28,3-29,3-30,3-31,3-32 on the "test item and command" sheet. (item number modification)
Added In-band Spurious (1M / 2M) and Out-band Spurious (1M / 2M) content to the "Test Specs" sheet.

## Measurement system
Measurement system
Test process outline
Ref
ANT
Shielding BOX
RF (2.4GHz)
USB
or
LAN
UART_RX
UART_TX
UART_TX
UART_RX
USB
USB-UART Converter
Socket and board
ANT
USB
VDD
Tester
DAQM901A
Current
DUT
(TSA8C1)
USB or GPIB or LAN
DC power supply
VDDIO
VDD_NRF
connector
USB or GPIB or LAN
connector
Tester
MT8872
GPIB
or
LAN
Test port B
Test port A
RF (2.4GHz)
RFIO
Frequency measurement
P0.20
SWDCLK
SWDIO
Programmer
(ex. Flasher ARM)
SWDCLK
SWDIO
USB
PC
SPDT
SMA
1.FW write
・Test FW write
2.RF check (OTA)
*SPDT connects between RFIO and ANT
・Antenna test
・Measurement of current consumption during RF operation
3.RF check (conductive)
* SPDT connects between RFIO and MT8872
・RF performance test
・Conformity test of Radio Law (construction design certification)
*Only "out-band Spurious Emissions" are measured using a spectrum analyzer.
4.Function check
・32kHz Xtal operation check
・DCDC operating current check
・Sleep current check
・Test FW erase
*SPDT:Shielding BOX IN or OUT
SMA
SMA
Spectrum
analyzer
XXXX
SPDT
port
*SPDT:Since it is used for spurious measurement, the corresponding
frequency range must be 12.5 GHz or higher.

## Command Format
1.1.1.Command Format |  |  |  | Command sequence
Request command ： Len Command ID Data( If necessary ） Respons command ： Len Command ID¦0x80 Result Code + Data( If necessary ）
1.2.Request command(PC→DUT)
Len | Command ID | Data
Command ID + Data length | Command ID(*1） | command Parameter
1.3.Response command（DUT→PC)
Len | Command ID | Result Code | Data
Command ID+Data length | Command（*1） | Result Code | Response Data
*1 Comand List
The response command ID is the logical OR of the request command ID and 0x80. Please refer to an example on the right side. |  |  |  |  | <Example>
Write an OK flag in the FW Write process Flag area
Command： Write Process Flag
If it send Command: 0x03 01 00 00, it will get a response of Command Respons: 0x02 81 00.
1.3.1.Command List
1.3.1.1.Request command(PC→DUT) |  |  |  |  | <Example>
Command： Read FW Verison
Command | Command ID | Data |  |  | If it send Command: 0x01 06, it will get a response of Command Respons: 0x04 86 00 X1 X2.
Write Process Flag | 0x01 | Data size is 2Byte
Flag area(1byte) + Write Flag Value(1byte)
<Flag area>
FW Write process Flag area :0x00
OTA process Flag area :0x01
Conductive process Flag area :0x02
Function process Flag area :0x03
<Write Flag Value>
OK Flag:0x00
NG Flag:0x01
Read Process Flag | 0x02 | Data size is 1Byte
Flag area(1byte)
< Flag area >
FW Write:0x00
OTA:0x01
Conductive:0x02
Function:0x03
All area:0xFF
Set DTM | 0x03 | Data size is 1byte
Power mode(1byte)
<Power mode>
LDO mode: 0x00
DCDC mode: 0x01
End DTM | 0xFF | FF (The request command is fixed to "0xFF FF" because it is a command used in DTM.)
Read BD Address | 0x05 | -
Read FW Ver | 0x06 | -
Set Sleep Mode | 0x07 | -
Check 32.768kHz Xtal FREQ | 0x08 | Data size is 1byte
Enable flag(1byte)
<Enable flag>
Output disable: 0x00
Output enable: 0x01
SET DCDC | 0x09 | Data size is 1byte
Enable flag(1byte)
<Enable flag>
DCDC disable: 0x00
DCDC enable: 0x01
Init Settings | 0x0A | -
1.3.1.2.Response command(DUT→PC)
Command | Command ID | Result Code | Data
Write Process Flag | 0x81 | Success:0x00 | -
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
"Flag area" is an invalid value: 0x01
Failed to write FLASH: 0x02
Read Process Flag | 0x82 | Success:0x00 | Data size is 2Byte or 5Byte
Flag area(1byte) ＋ Flag Value(1byte or 4byte))
< Flag area >
FW Write:0x00
OTA:0x01
Conductive:0x02
Function:0x03
All:0xFF
<Flag Value>
OK Flag:0x00
NG Flag:0x01
Ex) IF Flag area 0x00～0x03
Flag area(1byte) + Flag Value(1byte)
Ex) IF Flag area 0xFF
Flag area(1byte) + All Area Flag Value(4byte)
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
"Flag area" is an invalid value: 0x01
Set DTM | 0x83 | Success:0x00 | -
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
"Power mode" is an invalid value: 0x01
End DTM | 0x84 | Success:0x00 | -
Read BD Address | 0x85 | Success:0x00 | Data size is 6byte.
BD Address
0xX1 X2 X3 X4 X5 X6
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
Read FW Ver | 0x86 | Success:0x00 | Data size is 2byte.
FW Version
0xX1 X2 (TBD)
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
Set Sleep Mode | 0x87 | Success:0x00 | -
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
Check 32.768kHz Xtal FREQ | 0x88 | Success:0x00 | -
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
"Enable flag" is an invalid value: 0x01
SET DCDC | 0x89 | Success:0x00 | -
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
"Enable flag" is an invalid value: 0x01
Init settings | 0x8A | Success:0x00 | -
|  | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" does not match the command format: 0x00
*An error that occurred while receiving the command.
- | 0xFF | Error :0x01 | Data size is 1Byte
Error code(1byte)
<Error code>
"Len" is an invalid value.(= 0x00 or 0xFF) : 0x00
Receive timeout: 0x01
"Command ID" is an invalid value: 0x02
PC
DUT
Request Command
Response Command

## control flow and sequence
FW Write Proccess | OTA Test Process | Coductive Test Process | Function Test Process
PC/Handler
DUT
Command : Check FW Version (ID=0x06)
Command Response : FW Version(ID=0x86)
Write “DUT FW”
UART connection
Reboot (HW Reset or Power off/on)
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Command : Read BD Address (ID=0x05)
Command Response : Read BD Address (ID=0x85)
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
All Process Flag init
Write “NG Valu” to Write Process Flag
DUT Initialize
Cheack DUT FW Version
Read DUT BD Address
Output Power Test(1Mbps)
Output Power Test(1Mbps)
ALL OK?
Output Power Test(2Mbps)
Output Power Test(2Mbps)
ALL OK?
Start
Write NG Flag in conductive process area
Set DTM
Go to End
Carrier Frequency Offset and Drift Test(1Mbps)
Carrier Frequency Offset and Drift Test(1Mbps) ALL OK?
Carrier Frequency Offset and Drift Test(2Mbps)
Carrier Frequency Offset and Drift Test(2Mbps) ALL OK?
Modulation Index Test(1Mbps)
Modulation Index Test(1Mbps)
ALL OK?
Modulation Index Test(2Mbps)
Modulation Index Test(2Mbps)
ALL OK?
Go to End
32.768kHz Xtal FREQ Test
32.768kHz Xtal FREQ Test
OK?
DCDC Check
DCDC Check OK?
Start
Write NG Flag in Function process area
Go to End
Command : Read Process Flag (ID=0x02)
Command Response : Read Process Flag (ID=0x82)
Erase “DUT FW”
DUT Power OFF
All process Tests are OK.
Power OFF
Case of Last Procces
Sleep Current Check
Sleep Current Check OK?
Reboot (HW Reset or Power off/on)
Write OK Flag in Function process area
DUT Power Off
End
All proccess Tests OK?
DUT FW Erase
If last process?
Go to End
Start
Go to next process
Set the DUT in the socket , DUT power on
Write NG Flag in FW writing process area
Connect the DUT to a Debugger / PC application
Write DUT FW at "J-Flash"
Reboot (HW Reset or Power off/on)
Connect DUT and test PC by UART
Read BD ADDR
Write OK Flag in FW writing process area
Check FW Version
All Proccess Flag initialize
Go to End
PC/Handler
DUT
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
ANT Test
TX Current Test
RX Current Test
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Command : Set DTM (ID=0x01)
Command Response : Set DTM Mode (ID=0x83)
Write “NG Valu” to OTA Process Flag
Set DTM Mode
Write “OK Valu” to OTA Process Flag
Command : End DTM (0xFF FF)
End DTM Mode
Command Response : End DTM Mode (ID=0x84)
Go ANT Test
ANT Test ALL OK?
Go TX Current Test
TX Current Test ALL OK?
RX Current Test
RX Current Test ALL OK?
Start
Write NG Flag in OTA process area
Set DTM
Write OK Flag in OTA process area
Go to End
Go to next process
End DTM
PC/Handler
DUT
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Command : Cheak 32.768kHz Xtal FREQ Output enable (ID=0x08)
Command Response : Cheak 32.768kHz Xtal FREQ (ID=0x88)
Command : SET DCDC enable (ID=0x09)
Command Response : SET DCDC (ID=0x89)
Command : Set Sleep Mode (ID=0x07)
Command Response : Set Sleep Mode (ID=0x87)
Reboot (HW Reset or Power off/on)
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Command : Check FW Version (ID=0x06)
Command Response : FW Version(ID=0x86)
Check DUT FW Version
Command : Cheak 32.768kHz Xtal FREQ Output disable (ID=0x08)
Command Response : Cheak 32.768kHz Xtal FREQ (ID=0x88)
Command : Init Settings (ID=0x0A)
Command Response : Init Settings (ID=0x8A)
PC/Handler
DUT
Command : Read BD Address (ID=0x05)
Command Response : Read BD Address (ID=0x85)
Command : Read Process Flag (ID=0x02)
Command Response : Read Process Flag (ID=0x82)
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Output Power Test(1Mbps)
Output Power Test(2Mbps)
Carrier Frequency Offset and Drift Test(1Mbps)
Command : Set DTM Mode (ID=0x01)
Command Response : Set DTM Mode (ID=0x81)
Carrier Frequency Offset and Drift Test(2Mbps)
Modulation Index Test(1Mbps)
Modulation Index Test(2Mbps)
Receiver Sensitivity Test(1Mbps)
Receiver Sensitivity Test(2Mbps)
Command : Write Process Flag (ID=0x01)
Command Response : Write Process Flag (ID=0x81)
Command : End DTM (0xFF FF)
End DTM Mode
Command Response : End DTM Mode (ID=0x84)
In-band Spurious Test(1M/2Mbps)
Out-band Spurious Test(1M/2Mbps)
In-band Spurious Test(1M/2Mbps)
In-band Spurious Test
ALL OK?
Out-band Spurious Test(1M/2Mbps)
Out-band Spurious Test
ALL OK?
Go to End
Receiver Sensitivity Test(1Mbps)
Receiver Sensitivity Test(1Mbps)
ALL OK?
Receiver Sensitivity Test(2Mbps)
Receiver Sensitivity Test(2Mbps)
ALL OK?
Write OK Flag in conductive process area
Go to next process
Go to End
End DTM

## test item and command
SASP8(TSA8C1) test item/command
[6 empty] | default:LDO | UART Baud rate: 115200 | [6 empty] | 試験ランク test rank
| item |  |  |  | Operation mode | Power mode (internal power regulator) | Interface | DUT supply voltage | Request command (PC→DUT) | Response command （DUT→PC) | Test spec (Factory spec) （TBD） | Test condition/comment | Output measurement results to log | 試作 Prototype | 量産 Mass production
| 1 | FW write | [11 empty] | ○ | ○
|  | 1 | DUT Power ON |  |  |  |  | 0V--->3.0V | [5 empty] | ○ | ○
|  | 2 | FW Write |  | DebugMode | LDO | SWD | 3.0V | - | - |  | Test FW write＠Use "J-Flash" After writing FW, disconnect from "J-Flash". |  | ○ | ○ | Rve0.2:Add comment
|  | 3 | Reboot |  | APP | LDO | UART(2-wire) | 3.0V | - | - |  | HW Reset(P0.21/REST=Low) or Power off/on |  | ○ | ○
|  | 4 | FW Ver Check |  | APP | LDO | UART(2-wire) | 3.0V | 0x01 06 | Success:0x04 86 00 X1 X2 Error :Other than success | Success if FW Ver is the specified value | FW Ver confirmation + startup confirmation FW Version:0xX1 X2 (TBD) | ○ | ○ | ○
|  | 5 | Process flag writing (Initialize) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 00 5A 0x03 01 01 5A 0x03 01 02 5A 0x03 01 03 5A | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> FW Write Process Flag area :0x00 OTA process Flag area :0x01 Conductive process Flag area :0x02 Function process Flag area :0x03 |  | ○ | ○
|  | 6 | Process flag reading |  | APP | LDO | UART(2-wire) | 3.0V | 0x02 02 FF | Success:0x07 82 00 FF 5A 5A 5A 5A Error :Other than success | Check that all Flag areas are "5A". |  |  | ○ | ○
|  | 7 | Process flag writing (Write NG flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 00 01 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> FW Write Process Flag area :0x00 |  | ○ | ○
|  | 8 | Read BD Address |  | APP | LDO | UART(2-wire) | 3.0V | 0x01 05 | 0x08 85 00 X1 X2 X3 X4 X5 X6 | Success if BD Address can be read | Use BD address as control number. BD Address: 0xX1 X2 X3 X4 X5 X6 Data size is 6byte. | ○ | ○ | ○
|  | 9 | Process flag writing (Write OK flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 00 00 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> FW Write Process Flag area :0x00 |  | ○ | ○
| 2 | RF check (OTA) | [11 empty] | ○ | ○
|  | 1 | Process flag writing (Next process NG flag writing) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 01 01 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> OTA process Flag area :0x01 |  | ○ | ○
|  | 2 | Set DTM |  | APP--->DTM | DCDC | UART(2-wire) | 3.0V | 0x02 03 01 | Success:0x02 83 00 Error :0x02 83 01 |  | Operation mode chenge |  | ○ | ○ | Rve0.2: 0x01 03 ⇒ 0x02 03 01
|  | 3 | ANT_Test | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | RF test setting: 1M PHY, 2402MHz (0ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x80 94 | RF test setting: Success:0x00 00 Error :Other than success |  | Test with SASP antenna connected. Measure the transmission POWER with SASP antenna connected. |  | ○ | ー
[9 empty] | measurement: Output Power | - | upper limit: -30dBm lower limit: -60dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 4 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | RF test setting: 1M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x93 94 | RF test setting: Success:0x00 00 Error :Other than success |  | Test with SASP antenna connected. Measure the transmission POWER with SASP antenna connected. |  | ○ | ○
[9 empty] | measurement: Output Power | - | upper limit: -30dBm lower limit: -60dBm |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
|  | 5 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | RF test setting: 1M PHY, 2480MHz (39ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0xA7 94 | RF test setting: Success:0x00 00 Error :Other than success |  | Test with SASP antenna connected. Measure the transmission POWER with SASP antenna connected. |  | ○ | ー
[9 empty] | measurement: Output Power | - | upper limit: -30dBm lower limit: -60dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 6 | TX_Current | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Common with ”2-2 ANT_Test” RF test setting: 1M PHY, 2402MHz (0ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x80 94 | RF test setting: Success:0x00 00 Error :Other than success |  | ”2-2 ANT_Test” Perform current measurement at the same time as measurement. Test with SASP antenna connected. Check current during transmission. The pins to be measured are as follows. RF Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ー | Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 15ｍA lower limit: 5ｍA |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 7 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Common with ”2-3 ANT_Test” RF test setting: 1M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x93 94 | RF test setting: Success:0x00 00 Error :Other than success |  | ”2-3 ANT_Test” Perform current measurement at the same time as measurement. Test with SASP antenna connected. Check current during transmission. The pins to be measured are as follows. RF Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ○ | Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 15ｍA lower limit: 5ｍA |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
|  | 8 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Common with ”2-4 ANT_Test” RF test setting: 1M PHY, 2480MHz (39ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0xA7 94 | RF test setting: Success:0x00 00 Error :Other than success |  | ”2-4 ANT_Test” Perform current measurement at the same time as measurement. Test with SASP antenna connected. Check current during transmission. The pins to be measured are as follows. RF Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ー | Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 15ｍA lower limit: 5ｍA |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 9 | RX_Current | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | RF test setting: 1M PHY, 2402MHz (0ch) 0x02 04 0x03 00 0x40 00 | RF test setting: Success:0x00 00 Error :Other than success |  | Test with SASP antenna connected. Check current when receiving. The pins to be measured are as follows. RF Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ー | Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 15ｍA lower limit: 5ｍA |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 10 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | RF test setting: 1M PHY, 2440MHz (19ch) 0x02 04 0x03 00 0x53 00 | RF test setting: Success:0x00 00 Error :Other than success |  | Test with SASP antenna connected. Check current when receiving. The pins to be measured are as follows. RF Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ○ | Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 15ｍA lower limit: 5ｍA |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
|  | 11 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | RF test setting: 1M PHY, 2480MHz (39ch) 0x02 04 0x03 00 0x67 00 | RF test setting: Success:0x00 00 Error :Other than success |  | Test with SASP antenna connected. Check current when receiving. The pins to be measured are as follows. RF Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ー | Rve0.4:Add comment
[9 empty] | measurement:Current | - | upper limit: 15ｍA lower limit: 5ｍA |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 12 | End DTM |  | DTM--->APP | LDO | UART(2-wire) | 3.0V | 0xFF FF | Success:0x02 84 00 Error :0x02 84 01 |  | Operation mode chenge |  | ○ | ○ | Rve0.2: 0x01 04 ⇒ 0xFF FF
|  | 13 | Process flag writing (Write OK flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 01 00 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> OTA process Flag area :0x01 |  | ○ | ○
| 3 | RF check (conductive) | [11 empty] | ○ | ○
|  | 1 | Process flag writing (Write NG flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 02 01 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> Conductive process Flag area :0x02 |  | ○ | ○
|  | 2 | Set DTM |  | APP--->DTM | DCDC | UART(2-wire) | 3.0V | 0x02 03 01 | Success:0x02 83 00 Error :0x02 83 01 |  | Operation mode chenge |  | ○ | ○ | Rve0.2: 0x01 03 ⇒ 0x02 03 01
|  | 3 | Output Power(1Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 1M PHY, 2402MHz (0ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x80 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-01-C [Output power] |  | ○ | ー
[9 empty] | measurement: Average Power | - | upper limit: 5dBm lower limit: 0dBm |  | ○ | ○
[9 empty] | Peak to Average | - | upper limit: 3dB |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 4 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 1M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x93 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-01-C [Output power] |  | ○ | ○
[9 empty] | measurement: Average Power | - | upper limit: 5dBm lower limit: 0dBm |  | ○ | ○ | ○
[9 empty] | Peak to Average | - | upper limit: 3dB |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
|  | 5 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 1M PHY, 2480MHz (39ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0xA7 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-01-C [Output power] |  | ○ | ー
[9 empty] | measurement: Average Power | - | upper limit: 5dBm lower limit: 0dBm |  | ○ | ○
[9 empty] | Peak to Average | - | upper limit: 3dB |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 6 | Output Power(2Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 2M PHY, 2402MHz (0ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0x80 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-01-C [Output power] |  | ○ | ー
[9 empty] | measurement: Average Power | - | upper limit: 5dBm lower limit: 0dBm |  | ○ | ○
[9 empty] | Peak to Average | - | upper limit: 3dB |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 7 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 2M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0x93 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-01-C [Output power] |  | ○ | ー
[9 empty] | measurement: Average Power | - | upper limit: 5dBm lower limit: 0dBm |  | ○ | ○
[9 empty] | Peak to Average | - | upper limit: 3dB |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 8 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 2M PHY, 2480MHz (39ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0xA7 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-01-C [Output power] |  | ○ | ー
[9 empty] | measurement: Average Power | - | upper limit: 5dBm lower limit: 0dBm |  | ○ | ○
[9 empty] | Peak to Average | - | upper limit: 3dB |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 9 | Carrier Frequency Offset and Drift (1Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Carrier Frequency test setting: 1M PHY, 2402MHz (0ch), (+4dBm),37byte,10101010 0x02 04 0x01 00 0x80 96 | Carrier Frequency test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-06-C [Carrier frequency offset and drift, uncoded data at 1 Ms/s] Payload: 10101010, Length: 37 Bytes, 1Mbps |  | ○ | ー
[9 empty] | measurement: Average Frequency Offset | - | - |  | ○ | ○
[9 empty] | Max +ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Max -ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Initial Drift Rate | - | upper limit: 23 kHz |  | ○ | ○
[9 empty] | Drift Rate | - | upper limit: 20 kHz / 50 µs |  | ○ | ○
[9 empty] | Max Drift | - | upper limit: 50 kHz |  | ○ | ○
[9 empty] | Average Drift | - | - |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 10 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Carrier Frequency test setting: 1M PHY, 2440MHz (19ch), (+4dBm),37byte,10101010 0x02 04 0x01 00 0x93 96 | Carrier Frequency test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-06-C [Carrier frequency offset and drift, uncoded data at 1 Ms/s] Payload: 10101010, Length: 37 Bytes, 1Mbps |  | ○ | ○
[9 empty] | measurement: Average Frequency Offset | - | - |  | ○ | ○ | ○
[9 empty] | Max +ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○ | ○
[9 empty] | Max -ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○ | ○
[9 empty] | Initial Drift Rate | - | upper limit: 23 kHz |  | ○ | ○ | ○
[9 empty] | Drift Rate | - | upper limit: 20 kHz / 50 µs |  | ○ | ○ | ○
[9 empty] | Max Drift | - | upper limit: 50 kHz |  | ○ | ○ | ○
[9 empty] | Average Drift | - | - |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
|  | 11 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Carrier Frequency test setting: 1M PHY, 2480MHz (39ch), (+4dBm),37byte,10101010 0x02 04 0x01 00 0xA7 96 | Carrier Frequency test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-06-C [Carrier frequency offset and drift, uncoded data at 1 Ms/s] Payload: 10101010, Length: 37 Bytes, 1Mbps |  | ○ | ー
[9 empty] | measurement: Average Frequency Offset | - | - |  | ○ | ○
[9 empty] | Max +ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Max -ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Initial Drift Rate | - | upper limit: 23 kHz |  | ○ | ○
[9 empty] | Drift Rate | - | upper limit: 20 kHz / 50 µs |  | ○ | ○
[9 empty] | Max Drift | - | upper limit: 50 kHz |  | ○ | ○
[9 empty] | Average Drift | - | - |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 12 | Carrier Frequency Offset and Drift (2Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Carrier Frequency test setting: 2M PHY, 2402MHz (0ch), (+4dBm),37byte,10101010 0x02 08 0x01 00 0x80 96 | Carrier Frequency test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-12-C[Carrier frequency offset and drift at 2 Ms/s] Payload: 10101010, Length: 37 Bytes, 2Mbps |  | ○ | ー
[9 empty] | measurement: Average Frequency Offset | - | - |  | ○ | ○
[9 empty] | Max +ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Max -ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Initial Drift Rate | - | upper limit: 23 kHz |  | ○ | ○
[9 empty] | Drift Rate | - | upper limit: 20 kHz / 50 µs |  | ○ | ○
[9 empty] | Max Drift | - | upper limit: 50 kHz |  | ○ | ○
[9 empty] | Average Drift | - | - |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 13 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Carrier Frequency test setting: 2M PHY, 2440MHz (19ch), (+4dBm),37byte,10101010 0x02 08 0x01 00 0x93 96 | Carrier Frequency test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-12-C[Carrier frequency offset and drift at 2 Ms/s] Payload: 10101010, Length: 37 Bytes, 2Mbps |  | ○ | ー
[9 empty] | measurement: Average Frequency Offset | - | - |  | ○ | ○
[9 empty] | Max +ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Max -ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Initial Drift Rate | - | upper limit: 23 kHz |  | ○ | ○
[9 empty] | Drift Rate | - | upper limit: 20 kHz / 50 µs |  | ○ | ○
[9 empty] | Max Drift | - | upper limit: 50 kHz |  | ○ | ○
[9 empty] | Average Drift | - | - |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 14 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Carrier Frequency test setting: 2M PHY, 2480MHz (39ch), (+4dBm),37byte,10101010 0x02 08 0x01 00 0xA7 96 | Carrier Frequency test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-12-C[Carrier frequency offset and drift at 2 Ms/s] Payload: 10101010, Length: 37 Bytes, 2Mbps |  | ○ | ー
[9 empty] | measurement: Average Frequency Offset | - | - |  | ○ | ○
[9 empty] | Max +ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Max -ve Frequency Offset | - | upper limit: 50kHz |  | ○ | ○
[9 empty] | Initial Drift Rate | - | upper limit: 23 kHz |  | ○ | ○
[9 empty] | Drift Rate | - | upper limit: 20 kHz / 50 µs |  | ○ | ○
[9 empty] | Max Drift | - | upper limit: 50 kHz |  | ○ | ○
[9 empty] | Average Drift | - | - |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 15 | Modulation Index (1Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Modulation Index test setting: (Delta f1) 1M PHY, 2402MHz (0ch), (+4dBm),37byte,11110000 0x02 04 0x01 00 0x80 95 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-05-C [Modulation Characteristics, uncoded data at 1 Ms/s] Payload: 11110000 & 10101010, Length: 37 Bytes, 1Mbps |  | ○ | ー
[9 empty] | measurement: Delta f1 max | - | - |  | ○ | ○
[9 empty] | Delta f1 average | - | 225 kHz ≤ F1avg ≤ 275 kHz |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
[9 empty] | Modulation Index test setting: (Delta f2) 1M PHY, 2402MHz (0ch), (+4dBm),37byte,10101010 0x02 04 0x01 00 0x80 96 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | ”3-9 Carrier Frequency Offset and Drift (1Mbps) Low Channel 2402MHz” Perform Delta f2 measurement at the same time as measurement. |  | ○
[9 empty] | Delta f2 max | - | - |  | ○ | ○
[9 empty] | Delta f2 average | - | ≥ 185 kHz |  | ○ | ○
[9 empty] | Delta f2avg / Delta f1avg | - | ≥ 0.8 |  | ○ | ○
[9 empty] | Delta f2 max % pass rate |  | ≥99.9% |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 16 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Modulation Index test setting: (Delta f1) 1M PHY, 2440MHz (19ch), (+4dBm),37byte,11110000 0x02 04 0x01 00 0x93 95 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-05-C [Modulation Characteristics, uncoded data at 1 Ms/s] Payload: 11110000 & 10101010, Length: 37 Bytes, 1Mbps |  | ○ | ○
[9 empty] | measurement: Delta f1 max | - | - |  | ○ | ○ | ○
[9 empty] | Delta f1 average | - | 225 kHz ≤ F1avg ≤ 275 kHz |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
[9 empty] | Modulation Index test setting: (Delta f2) 1M PHY, 2440MHz (19ch), (+4dBm),37byte,10101010 0x02 04 0x01 00 0x93 96 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | ”3-10 Carrier Frequency Offset and Drift (1Mbps) Mid Channel 2440MHz” Perform Delta f2 measurement at the same time as measurement. |  | ○ | ○
[9 empty] | Delta f2 max | - | - |  | ○ | ○ | ○
[9 empty] | Delta f2 average | - | ≥ 185 kHz |  | ○ | ○ | ○
[9 empty] | Delta f2avg / Delta f1avg | - | ≥ 0.8 |  | ○ | ○ | ○
[9 empty] | Delta f2 max % pass rate |  | ≥99.9% |  | ○ | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○
|  | 17 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Modulation Index test setting: (Delta f1) 1M PHY, 2480MHz (39ch), (+4dBm),37byte,11110000 0x02 04 0x01 00 0xA7 95 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-05-C [Modulation Characteristics, uncoded data at 1 Ms/s] Payload: 11110000 & 10101010, Length: 37 Bytes, 1Mbps |  | ○ | ー
[9 empty] | measurement: Delta f1 max | - | - |  | ○ | ○
[9 empty] | Delta f1 average | - | 225 kHz ≤ F1avg ≤ 275 kHz |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
[9 empty] | Modulation Index test setting: (Delta f2) 1M PHY, 2480MHz (39ch), (+4dBm),37byte,10101010 0x02 04 0x01 00 0xA7 96 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | ”3-11 Carrier Frequency Offset and Drift (1Mbps) High Channel 2480MHz” Perform Delta f2 measurement at the same time as measurement. |  | ○
[9 empty] | Delta f2 max | - | - |  |  | ○
[9 empty] | Delta f2 average | - | ≥ 185 kHz |  |  | ○
[9 empty] | Delta f2avg / Delta f1avg | - | ≥ 0.8 |  |  | ○
[9 empty] | Delta f2 max % pass rate |  | ≥99.9% |  |  | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 18 | Modulation Index (2Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Modulation Index test setting: (Delta f1) 2M PHY, 2402MHz (0ch), (+4dBm),37byte,11110000 0x02 08 0x01 00 0x80 95 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-10-C [Modulation Characteristics, at 2 Ms/s] Payload: 11110000 & 10101010, Length: 37 Bytes, 2Mbps |  | ○ | ー
[9 empty] | measurement: Delta f1 max | - | - |  | ○ | ○
[9 empty] | Delta f1 average | - | 450 kHz ≤ F1avg ≤ 550 kHz |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
[9 empty] | Modulation Index test setting: (Delta f2) 2M PHY, 2402MHz (0ch), (+4dBm),37byte,10101010 0x02 08 0x01 00 0x80 96 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | ”3-12 Carrier Frequency Offset and Drift (2Mbps) Low Channel 2402MHz” Perform Delta f2 measurement at the same time as measurement. |  | ○
[9 empty] | Delta f2 max | - | - |  | ○ | ○
[9 empty] | Delta f2 average | - | ≥ 370 kHz |  | ○ | ○
[9 empty] | Delta f2avg / Delta f1avg | - | ≥ 0.8 |  | ○ | ○
[9 empty] | Delta f2 max % pass rate |  | ≥99.9% |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 19 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V | Modulation Index test setting: (Delta f1) 2M PHY, 2440MHz (19ch), (+4dBm),37byte,11110000 0x02 08 0x01 00 0x93 95 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-10-C [Modulation Characteristics, at 2 Ms/s] Payload: 11110000 & 10101010, Length: 37 Bytes, 2Mbps |  | ○ | ー
[9 empty] | measurement: Delta f1 max | - | - |  | ○ | ○
[9 empty] | Delta f1 average | - | 450 kHz ≤ F1avg ≤ 550 kHz |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
[9 empty] | Modulation Index test setting: (Delta f2) 2M PHY, 2440MHz (19ch), (+4dBm),37byte,10101010 0x02 08 0x01 00 0x93 96 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | ”3-13 Carrier Frequency Offset and Drift (2Mbps) Mid Channel 2440MHz” Perform Delta f2 measurement at the same time as measurement. |  | ○
[9 empty] | Delta f2 max | - | - |  | ○ | ○
[9 empty] | Delta f2 average | - | ≥ 370 kHz |  | ○ | ○
[9 empty] | Delta f2avg / Delta f1avg | - | ≥ 0.8 |  | ○ | ○
[9 empty] | Delta f2 max % pass rate |  | ≥99.9% |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 20 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V | Modulation Index test setting: (Delta f1) 2M PHY, 2480MHz (39ch), (+4dBm),37byte,11110000 0x02 08 0x01 00 0xA7 95 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-10-C [Modulation Characteristics, at 2 Ms/s] Payload: 11110000 & 10101010, Length: 37 Bytes, 2Mbps |  | ○ | ー
[9 empty] | measurement: Delta f1 max | - | - |  | ○ | ○
[9 empty] | Delta f1 average | - | 450 kHz ≤ F1avg ≤ 550 kHz |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
[9 empty] | Modulation Index test setting: (Delta f2) 2M PHY, 2480MHz (39ch), (+4dBm),37byte,10101010 0x02 08 0x01 00 0xA7 96 | Modulation Index test setting: Success:0x00 00 Error :Other than success |  | ”3-14 Carrier Frequency Offset and Drift (2Mbps) High Channel 2480MHz” Perform Delta f2 measurement at the same time as measurement. |  | ○
[9 empty] | Delta f2 max | - | - |  | ○ | ○
[9 empty] | Delta f2 average | - | ≥ 370 kHz |  | ○ | ○
[9 empty] | Delta f2avg / Delta f1avg | - | ≥ 0.8 |  | ○ | ○
[9 empty] | Delta f2 max % pass rate |  | ≥99.9% |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 21 | In-band Spurious Emissions (1Mbps) | Low Channel 2406MHz | DTM | DCDC | UART(2-wire) | 3.0V | In-band Spurious Emissions test setting: 1M PHY, 2406MHz (2ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x82 94 | In-band Spurious Emissions test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-03-C [In-band emissions, uncoded data at 1 Ms/s] |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | measurement: In-Band Em.: Center Freq -5 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -4 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -3 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -2 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq (2406 MHz) |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +2 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +3 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +4 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +5 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 22 |  | Mid Channel 2440MHz |  |  |  |  | In-band Spurious Emissions test setting: 1M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x93 94 | In-band Spurious Emissions test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | measurement: In-Band Em.: Center Freq -5 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -4 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -3 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -2 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq (2440 MHz) |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +2 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +3 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +4 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +5 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 23 |  | High Channel 2476MHz |  |  |  |  | In-band Spurious Emissions test setting: 1M PHY, 2476MHz (37ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0xA5 94 | In-band Spurious Emissions test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | measurement: In-Band Em.: Center Freq -5 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -4 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -3 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -2 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq (2476 MHz) |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +2 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +3 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +4 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +5 MHz |  | upper limit:-30dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 24 | In-band Spurious Emissions (2Mbps) | Low Channel 2406MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 2M PHY, 2406MHz (2ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0x82 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | RFPHY/TRM/BV-08-C [In-band emissions at 2 Ms/s] |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | measurement: In-Band Em.: Center Freq -5 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -4 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -3 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -2 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq (2406 MHz) |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +2 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +3 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +4 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +5 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 25 |  | Mid Channel 2440MHz |  |  |  |  | Output Power test setting: 2M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0x93 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | measurement: In-Band Em.: Center Freq -5 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -4 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -3 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -2 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq (2440 MHz) |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +2 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +3 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +4 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +5 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 26 |  | High Channel 2476MHz |  |  |  |  | Output Power test setting: 2M PHY, 2476MHz (37ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0xA5 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | measurement: In-Band Em.: Center Freq -5 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -4 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -3 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -2 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq -1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq (2476 MHz) |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +1 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +2 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +3 MHz |  | - |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +4 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | In-Band Em.: Center Freq +5 MHz |  | upper limit:-20dBm |  | ○ | ○
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○
|  | 27 | Out-band Spurious Emissions (1Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Out-band Spurious Emissions test setting: 1M PHY, 2402MHz (0ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x80 94 | Out-band Spurious Emissions test setting: Success:0x00 00 Error :Other than success |  | Use a spectrum analyzer * If the measuring instrument selected by ASE has standard settings for spurious emission measurement, use that. |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | spectrum analyzer settings FREQ:CENT 2402MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | measurement: 1hrm (fo×N=2402MHzｘ1): Spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 4804MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 2hrm (fo×N=2402MHzｘ2): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 7206MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 3hrm (fo×N=2402MHzｘ3): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 9608MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 4hrm (fo×N=2402MHzｘ4): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 12010MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 5hrm (fo×N=2402MHzｘ5): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 30MHz / STOP 1GHz RBW:100 kHz VBW:100 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:30MHz-1GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(30MHz-1GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 1GHz / STOP 2.387GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:1GHz-2.387GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(1GHz-2.387GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.387GHz / STOP 2.400GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.387GHz-2.400GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.387GHz-2.400GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4835GHz / STOP 2.4965GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4835GHz-2.4965GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4835GHz-2.4965GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4965GHz / STOP 12.5GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4965GHz-12.5GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4965GHz-12.5GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ー
|  | 28 |  | Mid Channel 2440MHz |  |  |  |  | Out-band Spurious Emissions test setting: 1M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0x93 94 | Out-band Spurious Emissions test setting: Success:0x00 00 Error :Other than success |  | Whether or not there is a test in mass production depends on the test results in the prototype. |  | ○ | ○ or ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | spectrum analyzer settings FREQ:CENT 2440MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | measurement: 1hrm (fo×N=2440MHzｘ1): Spurious amplitude value |  | - |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:CENT 4880MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | 2hrm (fo×N=2440MHzｘ2): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:CENT 7232MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | 3hrm (fo×N=2440MHzｘ3): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:CENT 9760MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | 4hrm (fo×N=2440MHzｘ4): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:CENT 12200MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | 5hrm (fo×N=2440MHzｘ5): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:START 30MHz / STOP 1GHz RBW:100 kHz VBW:100 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | Out-Band Em.:30MHz-1GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(30MHz-1GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:START 1GHz / STOP 2.387GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | Out-Band Em.:1GHz-2.387GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(1GHz-2.387GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:START 2.387GHz / STOP 2.400GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | Out-Band Em.:2.387GHz-2.400GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.387GHz-2.400GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ○ or ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4835GHz / STOP 2.4965GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | Out-Band Em.:2.4835GHz-2.4965GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4835GHz-2.4965GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ○ or ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ○ or ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4965GHz / STOP 12.5GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ○ or ー
[9 empty] | Out-Band Em.:2.4965GHz-12.5GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4965GHz-12.5GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ○ or ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ○ or ー
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ○ or ー
|  | 29 |  | High Channel 2480MHz |  |  |  |  | Out-band Spurious Emissions test setting: 1M PHY, 2480MHz (39ch), (+4dBm),37byte,PRBS9 0x02 04 0x01 00 0xA7 94 | Out-band Spurious Emissions test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | spectrum analyzer settings FREQ:CENT 2480MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | measurement: 1hrm (fo×N=2480MHzｘ1): Spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 4960MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 2hrm (fo×N=2480MHzｘ2): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 7440MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 3hrm (fo×N=2480MHzｘ3): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 9920MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 4hrm (fo×N=2480MHzｘ4): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 12400MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 5hrm (fo×N=2480MHzｘ5): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 30MHz / STOP 1GHz RBW:100 kHz VBW:100 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:30MHz-1GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(30MHz-1GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 1GHz / STOP 2.387GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:1GHz-2.387GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(1GHz-2.387GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.387GHz / STOP 2.400GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.387GHz-2.400GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.387GHz-2.400GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4835GHz / STOP 2.4965GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4835GHz-2.4965GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4835GHz-2.4965GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4965GHz / STOP 12.5GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4965GHz-12.5GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4965GHz-12.5GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ー
|  | 30 | Out-band Spurious Emissions(2Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V | Output Power test setting: 2M PHY, 2406MHz (2ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0x82 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  | Use a spectrum analyzer * If the measuring instrument selected by ASE has standard settings for spurious emission measurement, use that. |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | spectrum analyzer settings FREQ:CENT 2402MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | measurement: 1hrm (fo×N=2402MHzｘ1): Spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 4804MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 2hrm (fo×N=2402MHzｘ2): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 7206MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 3hrm (fo×N=2402MHzｘ3): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 9608MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 4hrm (fo×N=2402MHzｘ4): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 12010MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 5hrm (fo×N=2402MHzｘ5): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 30MHz / STOP 1GHz RBW:100 kHz VBW:100 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:30MHz-1GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(30MHz-1GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 1GHz / STOP 2.387GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:1GHz-2.387GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(1GHz-2.387GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.387GHz / STOP 2.400GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.387GHz-2.400GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.387GHz-2.400GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4835GHz / STOP 2.4965GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4835GHz-2.4965GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4835GHz-2.4965GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4965GHz / STOP 12.5GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4965GHz-12.5GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4965GHz-12.5GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ー
|  | 31 |  | Mid Channel 2440MHz |  |  |  |  | Output Power test setting: 2M PHY, 2440MHz (19ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0x93 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | spectrum analyzer settings FREQ:CENT 2440MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | measurement: 1hrm (fo×N=2440MHzｘ1) |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 4880MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 2hrm (fo×N=2440MHzｘ2): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 7232MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 3hrm (fo×N=2440MHzｘ3): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 9760MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 4hrm (fo×N=2440MHzｘ4): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 12200MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 5hrm (fo×N=2440MHzｘ5): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 30MHz / STOP 1GHz RBW:100 kHz VBW:100 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:30MHz-1GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(30MHz-1GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 1GHz / STOP 2.387GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:1GHz-2.387GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(1GHz-2.387GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.387GHz / STOP 2.400GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.387GHz-2.400GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.387GHz-2.400GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4835GHz / STOP 2.4965GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4835GHz-2.4965GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4835GHz-2.4965GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4965GHz / STOP 12.5GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4965GHz-12.5GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4965GHz-12.5GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ー
|  | 32 |  | High Channel 2480MHz |  |  |  |  | Output Power test setting: 2M PHY, 2476MHz (37ch), (+4dBm),37byte,PRBS9 0x02 08 0x01 00 0xA5 94 | Output Power test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ー | Rve0.3: Add test Rve0.4: Add command
[9 empty] | spectrum analyzer settings FREQ:CENT 2480MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | measurement: 1hrm (fo×N=2480MHzｘ1): Spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 4960MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 2hrm (fo×N=2480MHzｘ2): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 7440MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 3hrm (fo×N=2480MHzｘ3): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 9920MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 4hrm (fo×N=2480MHzｘ4): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:CENT 12400MHz FREQ:SPAN 1MHZz RBW:10 kHz VBW:10 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | 5hrm (fo×N=2440MHzｘ5): Spurious amplitude value |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 30MHz / STOP 1GHz RBW:100 kHz VBW:100 kHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:30MHz-1GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(30MHz-1GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 1GHz / STOP 2.387GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:1GHz-2.387GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(1GHz-2.387GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.387GHz / STOP 2.400GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.387GHz-2.400GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.387GHz-2.400GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4835GHz / STOP 2.4965GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4835GHz-2.4965GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4835GHz-2.4965GHz)) |  | upper limit:-16dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | spectrum analyzer settings FREQ:START 2.4965GHz / STOP 12.5GHz RBW:1 MHz VBW:1 MHz Detector:POSitive TRAC:TYPE MAXH | - | - |  |  | ○ | ー
[9 empty] | Out-Band Em.:2.4965GHz-12.5GHz Spurious amplitude value (maximum spurious amplitude value within the measurement frequency band(2.4965GHz-12.5GHz)) |  | upper limit:-26dBm |  | ○ | ○ | ー
[9 empty] | Frequency of maximum spurious amplitude value |  | - |  | ○ | ○ | ー
[9 empty] | test end: 0xC0 00 | test end: Success:0x80 00 Error :Other than success |  |  |  | ○ | ー
|  | 33 | Receiver Sensitivity (1Mbps) | Low Channel 2402MHz | DTM | DCDC | UART(2-wire) | 3.0V or 3.6V | MT8872 transmission settings (Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 1Mbps, Output Power:-80dBm ,Packets:1500) | - | - | RFPHY/RCV/BV-01-C [Receiver sensitivity, uncoded data at 1 Ms/s] Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 1Mbps, Output Power:-80dBm ,Packets:1500 *If you cannot set the Dirty Transmitter, you can use Non Dirty. Do not put the DUT in the receiving state until the MT8872 transmission setting is completed. |  | ○ | ー
[9 empty] | Receiver Sensitivity test setting: 1M PHY, 2402MHz (0ch),Standard 0x02 04 0x01 00 0x03 00 0x40 00 | Receiver Sensitivity test setting: Success:0x00 00 Error :Other than success |  |  |  | ○
[9 empty] | MT8872 transmission start (1500 packet) | - | - |  |  | ○
[9 empty] | transmission completed (1500 packet) | - | - |  |  | ○
[9 empty] | test end (RX): 0xC0 00 | test end (RX): Success:Range from 5DC(0%) to 40E(30.8%) 0x85 DC ・・・（Packet Error Rate:0%） 0x84 0E ・・・（Packet Error Rate:30.8%） Error :Other than success | 0x5DC=1500,0x40E=1038 Rx Packets:1500/1500 (Packet Error Rate:0%) Rx Packets:1038/1500 (Packet Error Rate:30.8%) |  |  | ○
[9 empty] | measurement: Packet Error Rate | - | upper limit: 30.8 % |  | ○ | ○
[9 empty] | Total Packets Received | - | - |  | ○ | ○
[9 empty] | Total Packets Tested | - | - |  | ○ | ○
|  | 34 |  | Mid Channel 2440MHz | DTM | DCDC | UART(2-wire) | 3.0V or 3.6V | MT8872 transmission settings (Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 1Mbps, Output Power:-80dBm ,Packets:1500) | - | - | RFPHY/RCV/BV-01-C [Receiver sensitivity, uncoded data at 1 Ms/s] Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 1Mbps, Output Power:-80dBm ,Packets:1500 *If you cannot set the Dirty Transmitter, you can use Non Dirty. Do not put the DUT in the receiving state until the MT8872 transmission setting is completed. |  | ○ | ○
[9 empty] | Receiver Sensitivity test setting: 1M PHY, 2440MHz (19ch),Standard 0x02 04 0x01 00 0x03 00 0x53 00 | Receiver Sensitivity test setting: Success:0x00 00 Error :Other than success |  |  |  | ○ | ○
[9 empty] | MT8872 transmission start (1500 packet) | - | - |  |  | ○ | ○
[9 empty] | transmission completed (1500 packet) | - | - |  |  | ○ | ○
[9 empty] | test end (RX): 0xC0 00 | test end (RX): Success:Range from 5DC(0%) to 40E(30.8%) 0x85 DC ・・・（Packet Error Rate:0%） 0x84 0E ・・・（Packet Error Rate:30.8%） Error :Other than success | 0x5DC=1500,0x40E=1038 Rx Packets:1500/1500 (Packet Error Rate:0%) Rx Packets:1038/1500 (Packet Error Rate:30.8%) |  |  | ○ | ○
[9 empty] | measurement: Packet Error Rate | - | upper limit: 30.8 % |  | ○ | ○ | ○
[9 empty] | Total Packets Received | - | - |  | ○ | ○ | ○
[9 empty] | Total Packets Tested | - | - |  | ○ | ○ | ○
|  | 35 |  | High Channel 2480MHz | DTM | DCDC | UART(2-wire) | 3.0V or 3.6V | MT8872 transmission settings (Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 1Mbps, Output Power:-80dBm ,Packets:1500) | - | - | RFPHY/RCV/BV-01-C [Receiver sensitivity, uncoded data at 1 Ms/s] Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 1Mbps, Output Power:-80dBm ,Packets:1500 *If you cannot set the Dirty Transmitter, you can use Non Dirty. Do not put the DUT in the receiving state until the MT8872 transmission setting is completed. |  | ○ | ー
[9 empty] | Receiver Sensitivity test setting: 1M PHY, 2480MHz (39ch),Standard 0x02 04 0x01 00 0x03 00 0x67 00 | Receiver Sensitivity test setting: Success:0x00 00 Error :Other than success |  |  |  | ○
[9 empty] | MT8872 transmission start (1500 packet) | - | - |  |  | ○
[9 empty] | transmission completed (1500 packet) | - | - |  |  | ○
[9 empty] | test end (RX): 0xC0 00 | test end (RX): Success:Range from 5DC(0%) to 40E(30.8%) 0x85 DC ・・・（Packet Error Rate:0%） 0x84 0E ・・・（Packet Error Rate:30.8%） Error :Other than success | 0x5DC=1500,0x40E=1038 Rx Packets:1500/1500 (Packet Error Rate:0%) Rx Packets:1038/1500 (Packet Error Rate:30.8%) |  |  | ○
[9 empty] | measurement: Packet Error Rate | - | upper limit: 30.8 % |  | ○ | ○
[9 empty] | Total Packets Received | - | - |  | ○ | ○
[9 empty] | Total Packets Tested | - | - |  | ○ | ○
|  | 36 | Receiver Sensitivity (2Mbps) | Low Channel：2402MHz | DTM | DCDC | UART(2-wire) | 3.0V or 3.6V | MT8872 transmission settings (Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 2Mbps, Output Power:-80dBm ,Packets:1500) | - | - | RF-PHY/RCV/BV-08-C [Receiver sensitivity at 2 Ms/s] Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 2Mbps, Output Power:-80dBm ,Packets:1500 *If you cannot set the Dirty Transmitter, you can use Non Dirty. Do not put the DUT in the receiving state until the MT8872 transmission setting is completed. |  | ○ | ー
[9 empty] | Receiver Sensitivity test setting: 2M PHY, 2402MHz (0ch),Standard 0x02 08 0x01 00 0x03 00 0x40 00 | Receiver Sensitivity test setting: Success:0x00 00 Error :Other than success |  |  |  | ○
[9 empty] | MT8872 transmission start (1500 packet) | - | - |  |  | ○
[9 empty] | transmission completed (1500 packet) | - | - |  |  | ○
[9 empty] | test end (RX): 0xC0 00 | test end (RX): Success:Range from 5DC(0%) to 40E(30.8%) 0x85 DC ・・・（Packet Error Rate:0%） 0x84 0E ・・・（Packet Error Rate:30.8%） Error :Other than success | 0x5DC=1500,0x40E=1038 Rx Packets:1500/1500 (Packet Error Rate:0%) Rx Packets:1038/1500 (Packet Error Rate:30.8%) |  |  | ○
[9 empty] | measurement: Packet Error Rate | - | upper limit: 30.8 % |  | ○ | ○
[9 empty] | Total Packets Received | - | - |  | ○ | ○
[9 empty] | Total Packets Tested | - | - |  | ○ | ○
|  | 37 |  | Mid Channel：2440MHz | DTM | DCDC | UART(2-wire) | 3.0V or 3.6V | MT8872 transmission settings (Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 2Mbps, Output Power:-80dBm ,Packets:1500) | - | - | RF-PHY/RCV/BV-08-C [Receiver sensitivity at 2 Ms/s] Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 2Mbps, Output Power:-80dBm ,Packets:1500 *If you cannot set the Dirty Transmitter, you can use Non Dirty. Do not put the DUT in the receiving state until the MT8872 transmission setting is completed. |  | ○ | ー
[9 empty] | Receiver Sensitivity test setting: 2M PHY, 2440MHz (19ch),Standard 0x02 08 0x01 00 0x03 00 0x53 00 | Receiver Sensitivity test setting: Success:0x00 00 Error :Other than success |  |  |  | ○
[9 empty] | MT8872 transmission start (1500 packet) | - | - |  |  | ○
[9 empty] | transmission completed (1500 packet) | - | - |  |  | ○
[9 empty] | test end (RX): 0xC0 00 | test end (RX): Success:Range from 5DC(0%) to 40E(30.8%) 0x85 DC ・・・（Packet Error Rate:0%） 0x84 0E ・・・（Packet Error Rate:30.8%） Error :Other than success | 0x5DC=1500,0x40E=1038 Rx Packets:1500/1500 (Packet Error Rate:0%) Rx Packets:1038/1500 (Packet Error Rate:30.8%) |  |  | ○
[9 empty] | measurement: Packet Error Rate | - | upper limit: 30.8 % |  | ○ | ○
[9 empty] | Total Packets Received | - | - |  | ○ | ○
[9 empty] | Total Packets Tested | - | - |  | ○ | ○
|  | 38 |  | High Channel：2480MHz | DTM | DCDC | UART(2-wire) | 3.0V or 3.6V | MT8872 transmission settings (Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 2Mbps, Output Power:-80dBm ,Packets:1500) | - | - | RF-PHY/RCV/BV-08-C [Receiver sensitivity at 2 Ms/s] Payload: PRBS 9, Length: 37 Bytes, Dirty Transmitter: 2Mbps, Output Power:-80dBm ,Packets:1500 *If you cannot set the Dirty Transmitter, you can use Non Dirty. Do not put the DUT in the receiving state until the MT8872 transmission setting is completed. |  | ○ | ー
[9 empty] | Receiver Sensitivity test setting: 2M PHY, 2480MHz (39ch),Standard 0x02 08 0x01 00 0x03 00 0x67 00 | Receiver Sensitivity test setting: Success:0x00 00 Error :Other than success |  |  |  | ○
[9 empty] | MT8872 transmission start (1500 packet) | - | - |  |  | ○
[9 empty] | transmission completed (1500 packet) | - | - |  |  | ○
[9 empty] | test end (RX): 0xC0 00 | test end (RX): Success:Range from 5DC(0%) to 40E(30.8%) 0x85 DC ・・・（Packet Error Rate:0%） 0x84 0E ・・・（Packet Error Rate:30.8%） Error :Other than success | 0x5DC=1500,0x40E=1038 Rx Packets:1500/1500 (Packet Error Rate:0%) Rx Packets:1038/1500 (Packet Error Rate:30.8%) |  |  | ○
[9 empty] | measurement: Packet Error Rate | - | upper limit: 30.8 % |  | ○ | ○
[9 empty] | Total Packets Received | - | - |  | ○ | ○
[9 empty] | Total Packets Tested | - | - |  | ○ | ○
|  | 39 | End DTM |  | DTM--->APP | LDO | UART(2-wire) | 3.0V | 0xFF FF | Success:0x02 84 00 Error :0x02 84 01 |  | Operation mode chenge |  | ○ | ○ | Rve0.2: 0x01 04 ⇒ 0xFF FF
|  | 40 | Process flag writing (Write OK flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 02 00 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> Conductive process Flag area :0x02 |  | ○ | ○
| 4 | Function check | [11 empty] | ○ | ○
|  | 1 | Process flag writing (Write NG flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 03 01 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> Conductive process Flag area :0x03 |  | ○ | ○
|  | 2 | 32.768kHz Xtal FREQ |  | APP | LDO | UART(2-wire) | 3.0V | Check 32.768kHz Xtal FREQ Output enable: 0x02 08 01 | Success:0x02 88 00 Error :0x02 88 01 |  | Output to P0.20 Measure 16.384kHz±50PPM (32.768kHz/2＝16.384kHz) The pins to be measured are as follows. 32.768kHz Xtal FREQ：P0.20(K4)ーGND ()：TSA8C1 Pin number |  | ○ | ○ | Rve0.2: 0x01 08 ⇒ 0x02 08 01 Rve0.4:Add comment
[9 empty] | measurement: FREQ | - | upper limit: 16.3848192kHz lower limit: 16.3831808kHz |  | ○ | ○ | ○
[9 empty] | Check 32.768kHz Xtal FREQ Output disable: 0x02 08 00 | Success:0x02 88 00 Error :0x02 88 01 |  |  |  | ○ | ○ | Rve0.2: TBD ⇒ 0x02 08 00
|  | 3 | DCDC Check |  | APP | LDO⇒DCDC | UART(2-wire) | 3.0V | SET DCDC DCDC enable: 0x02 09 01 | Success:0x02 89 00 Error :0x02 89 01 |  | Active current check in DCDC mode The pins to be measured are as follows. DCDC Check：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ○ | Rve0.2: 0x01 09 ⇒ 0x02 09 01 Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 2.5ｍA lower limit: 0.5ｍA |  | ○ | ○ | ○
[6 empty] | DCDC⇒LDO |  |  | Init Settings: 0x01 0A | Success:0x02 8A 00 Error :0x02 8A 01 |  |  |  | ○ | ○ | Rve0.2: 0x01 0A
|  | - | GPIO Port Check |  | APP | DCDC | UART(2-wire) |  |  |  |  | Each GPIO High / Low check (16 GPIOs) |  | ○ | ○ | ←delete GPIO inspection is carried out by O / S Diode test in the assembly process.
|  | 4 | Sleep Current |  | APP | LDO | UART(2-wire) | 3.0V | 0x01 07 | Success:0x02 87 00 Error :0x02 87 01 |  | Measurement accuracy in the μA range is required. DeepSleep needs to be restarted after measurement. The pins to be measured are as follows. Sleep Current：VDD(D1)-GND ()：TSA8C1 Pin number |  | ○ | ○ | Rve0.4:Add comment
[9 empty] | measurement: Current | - | upper limit: 15uA lower limit: 0.1uA |  | ○ | ○ | ○
|  | 5 | Reboot |  | APP | LDO | UART(2-wire) | 3.0V | - | - |  | HW Reset(P0.21/REST=Low) or Power off/on |  | ○ | ○ | Rve0.2: "Power off/on" comment delete.
|  | 6 | FW Ver Check |  | APP | LDO | UART(2-wire) | 3.0V | 0x01 06 | Success:0x04 86 00 X1 X2 Error :Other than success | Success if FW Ver is the specified value | Confirmation of startup after Reboot FW Version:0xX1 X2 (TBD) |  | ○ | ○
|  | 7 | Process flag writing (Write OK flag) |  | APP | LDO | UART(2-wire) | 3.0V | 0x03 01 03 00 | Success:0x02 81 00 Error :0x02 81 01 |  | <Flag area> Function process Flag area :0x03 |  | ○ | ○
|  | 8 | Process flag reading |  | APP | LDO | UART(2-wire) | 3.0V | 0x02 02 FF | Success:0x07 82 00 FF 00 00 00 00 Error :Other than success | Check that all Flag areas are "00". |  |  | ○ | ○
|  | 9 | FW erase（All Clear） |  | DebugMode | LDO | SWD | 3.0V | - | - |  | Test FW erase＠Use "J-Flash" |  | ○ | ○
|  | 10 | DUT Power OFF |  |  |  |  | 3.0V--->0V | [5 empty] | ○ | ○
DebugMode:SWD Connection
APP：Application Mode
DTM：Direct Test Mode
DTM (Direct Test Mode)
The UART interface characteristics shall be set to use the following parameters:
■Baud rate: One of the following shall be supported by the DUT:
1200, 2400, 9600, 14400, 19200, 38400, 57600, 115200
■Number of data bits: 8
■No parity
■1 stop bit
■No flow control (RTS or CTS)
See Bluetooth Core Specification (V5.2) "DIRECT TEST MODE" chapter for command specifications.
(3191page - 3209page)
About the sequence of test items:
It is okay to change the test items in the optimal sequence in ASE.

## Log output format
Log output item (draft)
Two types of logs are assumed.
■log
txt file (Format that can be converted to CSV file)
Model name | model Rev | Lot No | DATE/TIME | Measuring instrument ID | Test line No | Program version (FW for PC) | Program Version (FW for DUT) | BD Address | Test AA | Test AA (measured value) | ・・・ | ・・・ | Test XX | Test XX (measured value)
TSA8C1 | 01 | [6 empty] | XX:XX:XX:XX:XX:XX | OK／NG | XXXX |  |  | OK／NG | XXXX
TSA8C1 | 01 | [6 empty] | XX:XX:XX:XX:XX:XX | OK／NG | XXXX |  |  | OK／NG | XXXX
TSA8C1 | 01 | [6 empty] | XX:XX:XX:XX:XX:XX | OK／NG | XXXX |  |  | OK／NG | XXXX
・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・
・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・
・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・ | ・
■Individual log （ Individual Log for each DUT ）
txt file | Prototype only (can be changed with or without output)
Time stamp | Send command
Time stamp | Receive command
Time stamp | Send command
Time stamp | Receive command
Time stamp | Send command
Time stamp | Receive command
Time stamp | measured value
Time stamp | Send command
Time stamp | Receive command
Time stamp | measured value
・ | ・
・ | ・
・ | ・

## Test spec
Environmental condition
| No. | Item | conditions |  |  | note
| 1 | Temperature | ＋20 ～ +35℃ (TBD)
| 2 | Humidity | 20 ～ 75%RH (TBD)
| 3 | Power supply voltage | 1.7V±1%／3.0V±1%／3.6V±1% (TBD)
Test spec (TBD)
| No. | Item | unit | Test spec (Factory spec) | Design median | note
| 2 | RF check (OTA) |  |  |  | VDD=3.0V
ANT Test
| 2-3 | Low Channel(2402MHz) | dBm | upper limit:-60 lower limit:-30 | ｰ45 | Transmission Power confirmation including antenna TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, 1Mbps
| 2-4 | Mid Channel(2440MHz) | dBm | upper limit:-60 lower limit:-30 | ｰ45
| 2-5 | High Channel(2480MHz) | dBm | upper limit:-60 lower limit:-30 | ｰ45
TX Current
| 2-6 | Low Channel(2402MHz) | mA | 5 ≦ Current ≦ 15 | 8 | Average current during TX operation TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:1M
| 2-7 | Mid Channel(2440MHz) | mA | 5 ≦ Current ≦ 15 | 8
| 2-8 | High Channel(2480MHz) | mA | 5 ≦ Current ≦ 15 | 8
RX Current
| 2-9 | Low Channel(2402MHz) | mA | 5 ≦ Current ≦ 15 | 8 | Average current during RX operation PHY:1M
| 2-10 | Mid Channel(2440MHz) | mA | 5 ≦ Current ≦ 15 | 8
| 2-11 | High Channel(2480MHz) | mA | 5 ≦ Current ≦ 15 | 8
| No. | Item | unit | Test spec (Factory spec) | Design median | note
| 3 | RF check (conductive) |  |  |  | VDD=3.0V
TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:1M
Output Power(1Mbps)
| 3-3 | Low Channel(2402MHz) | dBm | 0 ≦ AvgPower ≦ 5 | 4
|  |  | dB | Peak to Average ≦ 3
| 3-4 | Mid Channel(2440MHz) | dBm | 0 ≦ AvgPower ≦ 5 | 4
|  |  | dB | Peak to Average ≦ 3
| 3-5 | High Channel(2480MHz) | dBm | 0 ≦ AvgPower ≦ 5 | 4
|  |  | dB | Peak to Average ≦ 3
|  | Output Power(2Mbps) |  |  |  | TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:2M
| 3-6 | Low Channel(2402MHz) | dBm | 0 ≦ AvgPower ≦ 5 | 4
|  |  | dB | Peak to Average ≦ 3
| 3-7 | Mid Channel(2440MHz) | dBm | 0 ≦ AvgPower ≦ 5 | 4
|  |  | dB | Peak to Average ≦ 3
| 3-8 | High Channel(2480MHz) | dBm | 0 ≦ AvgPower ≦ 5 | 4
|  |  | dB | Peak to Average ≦ 3
Carrier Frequency Offset and Drift (1Mbps)
| 3-9 | Low Channel(2402MHz) | kHz | Max +ve Frequency Offset ≤ 50 kHz Max -ve Frequency Offset ≤ 50 kHz Initial Drift Rate < 23 kHz Drift Rate / 50 µs < 20 kHz / 50 µs Max Drift < 50 kHz |  | TX Power:+4dBm Payload: 10101010 Length: 37 Bytes PHY:1M
| 3-10 | Mid Channel(2440MHz) | kHz | Max +ve Frequency Offset ≤ 50 kHz Max -ve Frequency Offset ≤ 50 kHz Initial Drift Rate < 23 kHz Drift Rate / 50 µs < 20 kHz / 50 µs Max Drift < 50 kHz
| 3-11 | High Channel(2480MHz) | kHz | Max +ve Frequency Offset ≤ 50 kHz Max -ve Frequency Offset ≤ 50 kHz Initial Drift Rate < 23 kHz Drift Rate / 50 µs < 20 kHz / 50 µs Max Drift < 50 kHz
Carrier Frequency Offset and Drift (2Mbps)
| 3-12 | Low Channel(2402MHz) | kHz | Max +ve Frequency Offset ≤ 50 kHz Max -ve Frequency Offset ≤ 50 kHz Initial Drift Rate < 23 kHz Drift Rate / 50 µs < 20 kHz / 50 µs Max Drift < 50 kHz |  | TX Power:+4dBm Payload: 10101010 Length: 37 Bytes PHY:2M
| 3-13 | Mid Channel(2440MHz) | kHz | Max +ve Frequency Offset ≤ 50 kHz Max -ve Frequency Offset ≤ 50 kHz Initial Drift Rate < 23 kHz Drift Rate / 50 µs < 20 kHz / 50 µs Max Drift < 50 kHz
| 3-14 | High Channel(2480MHz) | kHz | Max +ve Frequency Offset ≤ 50 kHz Max -ve Frequency Offset ≤ 50 kHz Initial Drift Rate < 23 kHz Drift Rate / 50 µs < 20 kHz / 50 µs Max Drift < 50 kHz
|  | Modulation Index (1Mbps) |  |  |  | TX Power:+4dBm Payload: 11110000 & 10101010 Length: 37 Bytes PHY:1M
| 3-15 | Low Channel(2402MHz) | kHz | 225 kHz ≤ F1avg ≤ 275 kHz F2avg ≥ 185 kHz f2avg / f1avg ≥ 0.8
|  |  | % | f2 max pass rate ≥ 99.9%
| 3-16 | Mid Channel(2440MHz) | kHz | 225 kHz ≤ F1avg ≤ 275 kHz F2avg ≥ 185 kHz f2avg / f1avg ≥ 0.8
|  |  | % | f2 max pass rate ≥ 99.9%
| 3-17 | High Channel(2480MHz) | kHz | 225 kHz ≤ F1avg ≤ 275 kHz F2avg ≥ 185 kHz f2avg / f1avg ≥ 0.8
|  |  | % | f2 max pass rate ≥ 99.9%
|  | Modulation Index (2Mbps) |  |  |  | 条件： TX Power:+4dBm Payload: 11110000 & 10101010 Length: 37 Bytes PHY:2M
| 3-18 | Low Channel(2402MHz) | kHz | 450 kHz ≤ F1avg ≤ 550 kHz F2avg ≥ 370 kHz f2avg / f1avg ≥ 0.8
|  |  | % | f2 max pass rate ≥ 99.9%
| 3-19 | Mid Channel(2440MHz) | kHz | 450 kHz ≤ F1avg ≤ 550 kHz F2avg ≥ 370 kHz f2avg / f1avg ≥ 0.8
|  |  | % | f2 max pass rate ≥ 99.9%
| 3-20 | High Channel(2480MHz) | kHz | 450 kHz ≤ F1avg ≤ 550 kHz F2avg ≥ 370 kHz f2avg / f1avg ≥ 0.8
|  |  | % | f2 max pass rate ≥ 99.9%
In-band Spurious Emissions(1Mbps)
| 3-21 | Low Channel(2406MHz) | dBm | -5MHz: upper limit:-30dBm -4MHz: upper limit:-30dBm -3MHz: upper limit:-30dBm -2MHz: upper limit:-20dBm |  | TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:1M
|  |  | dBm | +2MHz: upper limit:-20dBm +3MHz: upper limit:-30dBm +4MHz: upper limit:-30dBm +5MHz: upper limit:-30dBm
| 3-22 | Mid Channel(2440MHz) | dBm | -5MHz: upper limit:-30dBm -4MHz: upper limit:-30dBm -3MHz: upper limit:-30dBm -2MHz: upper limit:-20dBm
|  |  | dBm | +2MHz: upper limit:-20dBm +3MHz: upper limit:-30dBm +4MHz: upper limit:-30dBm +5MHz: upper limit:-30dBm
| 3-23 | High Channel(2476MHz) | dBm | -5MHz: upper limit:-30dBm -4MHz: upper limit:-30dBm -3MHz: upper limit:-30dBm -2MHz: upper limit:-20dBm
|  |  | dBm | +2MHz: upper limit:-20dBm +3MHz: upper limit:-30dBm +4MHz: upper limit:-30dBm +5MHz: upper limit:-30dBm
In-band Spurious Emissions(2Mbps)
| 3-24 | Low Channel(2406MHz) | dBm | -5MHz: upper limit:-20dBm -4MHz: upper limit:-20dBm |  | TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:2M
|  |  | dBm | +4MHz: upper limit:-20dBm +5MHz: upper limit:-20dBm
| 3-25 | Mid Channel(2440MHz) | dBm | -5MHz: upper limit:-20dBm -4MHz: upper limit:-20dBm
|  |  | dBm | +4MHz: upper limit:-20dBm +5MHz: upper limit:-20dBm
| 3-26 | High Channel(2476MHz) | dBm | -5MHz: upper limit:-20dBm -4MHz: upper limit:-20dBm
|  |  | dBm | +4MHz: upper limit:-20dBm +5MHz: upper limit:-20dBm
Out-band Spurious Emissions(1Mbps)
| 3-27 | Low Channel(2402MHz) | dBm | 2hrm: upper limit:-26dBm 3hrm: upper limit:-26dBm 4hrm: upper limit:-26dBm 5hrm: upper limit:-26dBm |  | TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:1M | upper limit:-26dBm(TELEC) upper limit:-41.3dBm(FCC)
|  |  | dBm | 30MHz-1GHz: upper limit:-26dBm(2.5uW/MHz) 1GHz-2.387GHz: upper limit:-26dBm(2.5uW/MHz) 2.387GHz-2.400GHz: upper limit:-16dBm(25uW/MHz) 2.4835GHz-2.4965GHz: upper limit:-16dBm(25uW/MHz) 2.4965GHz-12.5GHz: upper limit:-26dBm(2.5uW/MHz)
| 3-28 | Mid Channel(2440MHz) | dBm | 2hrm: upper limit:-26dBm 3hrm: upper limit:-26dBm 4hrm: upper limit:-26dBm 5hrm: upper limit:-26dBm
|  |  | dBm | 30MHz-1GHz: upper limit:-26dBm(2.5uW/MHz) 1GHz-2.387GHz: upper limit:-26dBm(2.5uW/MHz) 2.387GHz-2.400GHz: upper limit:-16dBm(25uW/MHz) 2.4835GHz-2.4965GHz: upper limit:-16dBm(25uW/MHz) 2.4965GHz-12.5GHz: upper limit:-26dBm(2.5uW/MHz)
| 3-29 | High Channel(2480MHz) | dBm | 2hrm: upper limit:-26dBm 3hrm: upper limit:-26dBm 4hrm: upper limit:-26dBm 5hrm: upper limit:-26dBm
|  |  | dBm | 30MHz-1GHz: upper limit:-26dBm(2.5uW/MHz) 1GHz-2.387GHz: upper limit:-26dBm(2.5uW/MHz) 2.387GHz-2.400GHz: upper limit:-16dBm(25uW/MHz) 2.4835GHz-2.4965GHz: upper limit:-16dBm(25uW/MHz) 2.4965GHz-12.5GHz: upper limit:-26dBm(2.5uW/MHz)
Out-band Spurious Emissions(2Mbps)
| 3-30 | Low Channel(2402MHz) | dBm | 2hrm: upper limit:-26dBm 3hrm: upper limit:-26dBm 4hrm: upper limit:-26dBm 5hrm: upper limit:-26dBm |  | TX Power:+4dBm Payload: PRBS9 Length: 37 Bytes, PHY:2M | upper limit:-26dBm(TELEC) upper limit:-41.3dBm(FCC)
|  |  | dBm | 30MHz-1GHz: upper limit:-26dBm(2.5uW/MHz) 1GHz-2.387GHz: upper limit:-26dBm(2.5uW/MHz) 2.387GHz-2.400GHz: upper limit:-16dBm(25uW/MHz) 2.4835GHz-2.4965GHz: upper limit:-16dBm(25uW/MHz) 2.4965GHz-12.5GHz: upper limit:-26dBm(2.5uW/MHz)
| 3-31 | Mid Channel(2440MHz) | dBm | 2hrm: upper limit:-26dBm 3hrm: upper limit:-26dBm 4hrm: upper limit:-26dBm 5hrm: upper limit:-26dBm
|  |  | dBm | 30MHz-1GHz: upper limit:-26dBm(2.5uW/MHz) 1GHz-2.387GHz: upper limit:-26dBm(2.5uW/MHz) 2.387GHz-2.400GHz: upper limit:-16dBm(25uW/MHz) 2.4835GHz-2.4965GHz: upper limit:-16dBm(25uW/MHz) 2.4965GHz-12.5GHz: upper limit:-26dBm(2.5uW/MHz)
| 3-32 | High Channel(2480MHz) | dBm | 2hrm: upper limit:-26dBm 3hrm: upper limit:-26dBm 4hrm: upper limit:-26dBm 5hrm: upper limit:-26dBm
|  |  | dBm | 30MHz-1GHz: upper limit:-26dBm(2.5uW/MHz) 1GHz-2.387GHz: upper limit:-26dBm(2.5uW/MHz) 2.387GHz-2.400GHz: upper limit:-16dBm(25uW/MHz) 2.4835GHz-2.4965GHz: upper limit:-16dBm(25uW/MHz) 2.4965GHz-12.5GHz: upper limit:-26dBm(2.5uW/MHz)
Receiver Sensitivity (1Mbps)
| 3-33 | Low Channel(2402MHz) | % | ≤ 30.8 % |  | 条件： Payload: PRBS 9 Length: 37 Bytes Dirty Transmitter:ON(1Mbps) *1 Output Power:-80dBm (TBD)
| 3-34 | Mid Channel(2440MHz) | % | ≤ 30.8 %
| 3-35 | High Channel(2480MHz) | % | ≤ 30.8 %
Receiver Sensitivity (2Mbps)
| 3-36 | Low Channel(2402MHz) | % | ≤ 30.8 % |  | Payload: PRBS 9 Length: 37 Bytes Dirty Transmitter: ON(2Mbps) *1 Output Power:-80dBm (TBD)
| 3-37 | Mid Channel(2440MHz) | % | ≤ 30.8 %
| 3-38 | High Channel(2480MHz) | % | ≤ 30.8 %
*1:If you cannot set the Dirty Transmitter, you can use Non Dirty.
| No. | Item | unit | Test spec (Factory spec) | Design median | note
| 4 | Function check |  |  |  | VDD=3.0V
| 4-2 | 32.768kHz Xtal FREQ | kHz | 16.3831808 ≦ FREQ ≦ 16.3848192 | 32.768／2 | 32.768kHz Xtal Check (±50ppm)
| 4-3 | DCDC Check | ｍA | 0.5 ≦ Current ≦ 2.5 | 1.2 | Average current during DCDC operation
| 4-4 | Sleep Current | uA | 0.1 ≦ Current ≦ 15 | 0.5 | Average current during DeepSleep operation

## BLE measurement item for MT8872
BLE measurement item for MT8872
Bluetooth TX Measurement
Bluetooth RX Measurement
Item No
2-3,2-4,2-5
3-3,3-4,3-5
3-6,3-7,3-8
Item No
3-15,3-16,3-17
Item No
3-9,3-10,3-11
Item No
3-18,3-19,3-20
Item No
3-12,3-13,3-14
Item No
3-27,3-28,3-29
Item No
3-30,3-31,3-32
Item No
3-24,3-25,3-26
Item No
3-21,3-22,3-23